# VitalCSI: Detailed Architecture and Methodology

This document summarizes the architecture and methodology presented in the paper 'VitalCSI: Contactless Respiratory Rate Estimation Using Consumer-Grade Wi-Fi Channel State Information'.

## 1. System Objective

VitalCSI is a contactless respiratory-rate monitoring system that estimates breathing using Wi-Fi Channel State Information (CSI). It employs commodity hardware including a Wi-Fi access point and Raspberry Pi running Nexmon CSI firmware. The design combines classical signal processing, dimensionality reduction, frequency analysis, time-domain breath detection, and Kalman filtering to achieve accurate respiratory-rate estimation across 6–33 breaths/min.

## 2. Overall Architecture

Hardware Layer
• ASUS Wi-Fi Access Point (802.11ac, 80 MHz)
• Raspberry Pi (VitalCSI receiver)
• Raspi-TX and Raspi-RX generate network traffic
• Subject positioned between AP and receiver

Software Layer
CSI Capture → Sliding Window → PCA → Component Quality Index (CQI)
→ Band-pass Filtering → FFT RR Estimation + Breath Counting
→ Signal Quality Assessment → Multidimensional Kalman Filter
→ Final Respiratory Rate.

## 3. Hardware Architecture

The receiver intercepts Wi-Fi packets in monitor mode using Nexmon CSI firmware. "
"An ASUS AC86U router transmits Wi-Fi packets while two Raspberry Pi devices continuously exchange ping packets. "
"The VitalCSI node passively captures CSI without modifying the router firmware. "
"256 OFDM subcarriers at 80 MHz bandwidth are available for analysis.

## 4. Methodology

Step 1: CSI Acquisition
• Capture CSI magnitude from every received packet.
• Sampling rate ≈50 packets/s.
• Use 30-second windows with 1-second overlap.

Step 2: Magnitude Extraction
• Only CSI magnitude is processed because commodity phase measurements are noisy.

Step 3: PCA
• Convert 256 correlated subcarriers into orthogonal principal components.
• Evaluate all components using quality metrics and retain the best five.

Step 4: Component Quality Index
Three metrics are computed:
1. Variance Index – retain components explaining significant variance.
2. Signal-to-Noise Ratio – maximise spectral energy inside the physiological breathing band.
3. Spectral Peak Index – reject components dominated by low-frequency drift.
CQI = Variance × SNR × Spectral Peak.
The five highest-CQI components are used downstream.

Step 5: Filtering
Apply high-pass and low-pass IIR filters to retain only physiological breathing frequencies (approximately 5–35 brpm).

Step 6: Dual Respiratory Estimation
Method A: FFT spectral peak detection.
Method B: Breath counting using the Box Slope Sum Function (BSSF).
Each PCA component yields two RR estimates.

Step 7: Signal Quality
Signal Quality Index (SQI) estimates confidence for every component.

Step 8: Kalman Fusion
A multidimensional Kalman filter fuses all RR estimates while accounting for temporal continuity and confidence, producing one robust RR estimate each second.

## 5. Mathematical Foundation

Wireless propagation is modelled as:
Y = Hx + n
where H is the CSI channel matrix, x is the transmitted signal and n is additive noise.
CSI is represented for every transmit antenna, receive antenna and OFDM subcarrier, encoding amplitude and phase changes caused by respiration.

## 6. Experimental Methodology

• 15 healthy volunteers.
• Controlled indoor room.
• Subject approximately 1 m from Wi-Fi devices.
• 20-minute recordings.
• Guided breathing from 6 to 33 breaths/min in 3-brpm increments every 2 minutes.
• Ground truth collected using a nasal airflow pressure sensor and pulse oximeter.

## 7. Performance

The pipeline achieved strong agreement with the airflow reference:
• Respiratory range: 6–33 brpm
• Mean Absolute Error: ≈1.20 brpm
• R² ≈0.93
The study demonstrates accurate contactless respiratory monitoring using only commodity Wi-Fi hardware.

## 8. Strengths

• Commodity hardware
• No wearable sensors
• Adaptive PCA selection
• Multiple complementary RR estimators
• Kalman fusion improves robustness
• Open-source implementation suitable for research.

## 9. Adaptation for ESP32/SilentSense

The methodology can be adapted directly to ESP32 CSI:
ESP32 CSI → Windowing → PCA → CQI → Filtering → FFT/Breath Counting
→ Kalman Fusion → Respiration → Higher-level modules (sleep, fall detection, digital twin, LLM reasoning).